## claim_ledger
claim_id | source_id | source | claim_text | claim_type | gspr | verification_question | required_evidence | status | primary_source | fallback_source | sota_reverse_correction | sota_corrected_claim_text | sota_support_status | sota_correction_basis | allowed_wording_strength
C-01 | SRC-230 | IFU intended purpose | The Bubble Study System is used in conjunction with Disposable Masks and Disposable Contrast Injection Tubing Set for the preparation and injection of agitated saline during bubble study procedures for transthoracic echocardiographic contrast imaging or contrast-enhanced transcranial Doppler ultrasound imaging. The NIBP module monitors patient pulse and non-invasive blood pressure. | intended_purpose | GSPR 1, 6, 23.4 | Does the clinical evidence support the IFU-defined use of automated agitated-saline preparation/injection for bubble-study RLS/PFO assessment? | SOTA benchmarks for agitated-saline bubble study, c-TTE/c-TCD RLS/PFO detection literature, subject-device clinical/PMS/PMCF evidence where available. | to_be_verified | PubMed/CT.gov | equivalent_device | keep_with_sota_limits | The Bubble Study System is used in conjunction with Disposable Masks and Disposable Contrast Injection Tubing Set for the preparation and injection of agitated saline during bubble study procedures for transthoracic echocardiographic contrast imaging or contrast-enhanced transcranial Doppler ultrasound imaging. The NIBP module monitors patient pulse and non-invasive blood pressure. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-02 | SRC-230 | IFU safety information / RMF | Risks associated with agitated-saline preparation and injection, disposable tubing use, contrast-bubble administration, patient cooperation, and monitoring are acceptable when controlled by IFU warnings, contraindications, risk controls, and PMS/PMCF. | safety | GSPR 1, 2, 8, 23.4 | Are injection-related, contrast-bubble, tubing, monitoring, and contraindication risks within accepted SOTA levels and controlled by IFU/RMF/PMS? | AE/SAE literature for agitated-saline bubble studies, vigilance/recall searches, IFU contraindications/warnings, RMF, PMS/PMCF. | to_be_verified | clinical+PMS+vigilance | PubMed/CT.gov | keep_with_sota_limits | Risks associated with agitated-saline preparation and injection, disposable tubing use, contrast-bubble administration, patient cooperation, and monitoring are acceptable when controlled by IFU warnings, contraindications, risk controls, and PMS/PMCF. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-03 | SRC-230 | IFU performance description | The system supports reproducible preparation and injection of agitated saline to obtain clinically interpretable c-TTE or c-TCD bubble-study results under the IFU-defined use conditions. | performance | GSPR 1, 6, 15 | Are diagnostic success, opacification adequacy, shunt detection/grading, and workflow performance aligned with SOTA acceptance criteria? | SOTA endpoints for bubble-study diagnostic success, RLS/PFO detection concordance, grading reproducibility, procedure completion, and device performance/PMS evidence. | to_be_verified | PubMed/CT.gov | bench_test | keep_with_sota_limits | The system supports reproducible preparation and injection of agitated saline to obtain clinically interpretable c-TTE or c-TCD bubble-study results under the IFU-defined use conditions. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-PERF-01 | SRC-230 | IFU performance / SOTA benchmark | The system prepares and injects agitated saline in a controlled, repeatable manner to support bubble-study contrast generation. | performance | GSPR 1, 6, 15 | Does evidence support repeatable agitated-saline preparation/injection performance compared with accepted manual or automated bubble-study practice? | SOTA benchmark BM-04, verification/PMS data on preparation consistency, device performance records, and relevant clinical literature. | to_be_verified | PubMed/CT.gov | bench_test | keep_with_sota_limits | The system prepares and injects agitated saline in a controlled, repeatable manner to support bubble-study contrast generation. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-PERF-02 | SRC-230 | IFU performance / clinical literature | The system enables adequate contrast-bubble opacification and procedure completion for c-TTE and/or c-TCD bubble-study assessment. | performance | GSPR 1, 6, 15 | Are opacification adequacy and diagnostic/procedure success aligned with SOTA for agitated-saline bubble studies? | SOTA benchmarks BM-01/BM-05, clinical literature on c-TTE/c-TCD bubble-study success, and subject-device clinical/PMS data where available. | to_be_verified | PubMed/CT.gov | bench_test | keep_with_sota_limits | The system enables adequate contrast-bubble opacification and procedure completion for c-TTE and/or c-TCD bubble-study assessment. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-PERF-03 | SRC-230 | IFU performance / SOTA benchmark | RLS/PFO detection and semi-quantitative shunt grading using the IFU-defined workflow are clinically interpretable against accepted bubble-study diagnostic pathways. | performance | GSPR 1, 6 | Are detection concordance, false positive/negative risk, and shunt grading reproducibility acceptable versus SOTA comparators? | SOTA benchmarks BM-02/BM-03/BM-07, guideline/pathway literature, and clinical validation or PMCF data. | to_be_verified | PubMed/CT.gov | bench_test | keep_with_sota_limits | RLS/PFO detection and semi-quantitative shunt grading using the IFU-defined workflow are clinically interpretable against accepted bubble-study diagnostic pathways. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-SAFE-01 | SRC-230 | IFU warnings / RMF | Injection-related adverse events, microbubble administration risks, and patient monitoring risks are controlled to acceptable levels through IFU instructions and risk controls. | safety | GSPR 1, 2, 8 | Are injection-related and cardiopulmonary/neurological safety events within SOTA accepted levels after IFU/RMF controls? | SOTA benchmark BM-06, AE/SAE literature for agitated-saline bubble studies, vigilance data, IFU contraindications/warnings, and RMF. | to_be_verified | clinical+PMS+vigilance | PubMed/CT.gov | keep_with_sota_limits | Injection-related adverse events, microbubble administration risks, and patient monitoring risks are controlled to acceptable levels through IFU instructions and risk controls. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-SAFE-02 | SRC-230 | IFU single-use tubing / sterility controls | Single-use contrast injection tubing and related disposable components control contamination, reuse, and performance-degradation risks when used as instructed. | safety | GSPR 2, 8, 11, 23.4 | Are tubing-set sterility, reuse-prevention, injection-path, and device-malfunction risks adequately controlled by source evidence? | Sterility/packaging evidence, RMF hazards, IFU warnings, PMS/PMCF complaints, and vigilance/recall searches. | to_be_verified | clinical+PMS+vigilance | PubMed/CT.gov | keep_with_sota_limits | Single-use contrast injection tubing and related disposable components control contamination, reuse, and performance-degradation risks when used as instructed. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-BEN-01 | SRC-230 | IFU intended purpose / clinical pathway | The system supports clinical assessment of suspected right-to-left shunt or patent foramen ovale by standardizing agitated-saline bubble-study workflow. | clinical_benefit | GSPR 1, 6, 23.4 | Does evidence support that standardized bubble-study workflow contributes to clinically useful RLS/PFO assessment? | Clinical pathway literature, SOTA benchmarks BM-01/BM-02/BM-03, subject-device clinical data, PMS/PMCF plan/data. | to_be_verified | PubMed/CT.gov | equivalent_device | keep_with_sota_limits | The system supports clinical assessment of suspected right-to-left shunt or patent foramen ovale by standardizing agitated-saline bubble-study workflow. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-BEN-02 | SRC-230 | IFU workflow / SOTA benchmark | Automated or controlled agitated-saline preparation may reduce manual variability, with the clinical conclusion limited to the strength of available evidence. | clinical_benefit | GSPR 1, 6 | Is reduced manual variability supported by direct performance, usability, or clinical workflow evidence? | Preparation consistency data, usability/workflow evidence, SOTA benchmark BM-04/BM-05, PMCF objectives if direct evidence is limited. | to_be_verified | PubMed/CT.gov | equivalent_device | keep_with_sota_limits | Automated or controlled agitated-saline preparation may reduce manual variability, with the clinical conclusion limited to the strength of available evidence. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-04 | SRC-230 | IFU warnings/contraindications | Patients with severe cardiac, hepatic, renal, or pulmonary impairment; malignant tumors; hematological or autoimmune diseases; | warning_contraindication | GSPR 1, 8, 23.4 | Is this warning/contraindication traceable to clinical risk evidence and risk controls? | IFU, RMF, vigilance, clinical literature. | to_be_verified | RMF/GSPR | PMS | keep_with_sota_limits | Patients with severe cardiac, hepatic, renal, or pulmonary impairment; malignant tumors; hematological or autoimmune diseases; (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-05 | SRC-230 | IFU warnings/contraindications | Patients unable to cooperate with the examination due to consciousness or cognitive disorders; | warning_contraindication | GSPR 1, 8, 23.4 | Is this warning/contraindication traceable to clinical risk evidence and risk controls? | IFU, RMF, vigilance, clinical literature. | to_be_verified | RMF/GSPR | PMS | keep_with_sota_limits | Patients unable to cooperate with the examination due to consciousness or cognitive disorders; (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-06 | SRC-230 | IFU warnings/contraindications | Patients with carotid artery stenosis, congenital intracranial segment vertebral artery (VA) dysplasia, or those without a temporal window and unable to locate the unilateral middle cerebral artery (MCA); | warning_contraindication | GSPR 1, 8, 23.4 | Is this warning/contraindication traceable to clinical risk evidence and risk controls? | IFU, RMF, vigilance, clinical literature. | to_be_verified | RMF/GSPR | PMS | keep_with_sota_limits | Patients with carotid artery stenosis, congenital intracranial segment vertebral artery (VA) dysplasia, or those without a temporal window and unable to locate the unilateral middle cerebral artery (MCA); (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits
C-07 | SRC-230 | IFU warnings/contraindications | Pregnant or lactating women. | warning_contraindication | GSPR 1, 8, 23.4 | Is this warning/contraindication traceable to clinical risk evidence and risk controls? | IFU, RMF, vigilance, clinical literature. | to_be_verified | RMF/GSPR | PMS | keep_with_sota_limits | Pregnant or lactating women. (supported within the limits of the SOTA benchmark and evidence applicability). | supported_by_sota_context | SOTA reverse correction constrains claim wording according to available benchmark value, source, population and evidence strength. | supported within SOTA limits